## Slide 1: Basic Crystallography – Data collection and reduction
Louise N. Dawe, PhD
Wilfrid Laurier University
Department of Chemistry and Biochemistry

## Slide 2: Basic Crystallography – Data collection and reduction
Louise N. Dawe, PhD
Wilfrid Laurier University
Department of Chemistry and Biochemistry
Narrated by Brian Patrick, PhD
University of British Columbia
Department of Chemistry

## Slide 3
Faculty of Science, Bijvoet Center for Biomolecular Research, Crystal and Structural Chemistry. ‘Interpretation of Crystal Structure Determinations’ 2005 Course Notes: http://www.cryst.chem.uu.nl/huub/notesweb.pdf
The University of Oklahoma: Chemical Crystallography Lab. Crystallography Notes and Manuals. http://xrayweb.chem.ou.edu/notes/index.html
Müller, P. Crystallographic Reviews, 2009, 15(1), 57-83.
Müller, Peter. 5.069 Crystal Structure Analysis, Spring 2010. (Massachusetts Institute of Technology: MIT OpenCourseWare), http://ocw.mit.edu/courses/chemistry/5-069-crystal-structure-analysis-spring-2010/. License: Creative Commons BY-NC-SA
References and Additional Resources

## Slide 4
X-ray Crystallography
Notes: 4

## Slide 5
Data Collection and Processing
Select and mount the crystal.
Center the crystal to the center of the goniometer circles (instrument maintenance.)
Collect several images; index the diffraction spots; refine the cell parameters; check for higher metric symmetry
Determine data collection strategy; collect data.
Reduce the data by applying background, profile (spot-shape), Lorentz, polarization and scaling corrections.
Determine precise cell parameters.
Collect appropriate information for an absorption correction. (Index the faces of the crystal. A highly redundant set of data is sufficient for an empirical absorption correction.)
Apply an absorption correction to the data.
(http://xrayweb.chem.ou.edu/notes/collect.html)

## Slide 6
Data Collection and Processing
Select and mount the crystal.
Center the crystal to the center of the goniometer circles (instrument maintenance.)
Collect several images; index the diffraction spots; refine the cell parameters; check for higher metric symmetry
Determine data collection strategy; collect data.
Reduce the data by applying background, profile (spot-shape), Lorentz, polarization and scaling corrections.
Determine precise cell parameters.
Collect appropriate information for an absorption correction. (Index the faces of the crystal. A highly redundant set of data is sufficient for an empirical absorption correction.)
Apply an absorption correction to the data.
(http://xrayweb.chem.ou.edu/notes/collect.html)

## Slide 7
Data Collection and Processing
Select and mount the crystal.
Center the crystal to the center of the goniometer circles (instrument maintenance.)
Collect several images; index the diffraction spots; refine the cell parameters; check for higher metric symmetry
Determine data collection strategy; collect data.
Reduce the data by applying background, profile (spot-shape), Lorentz, polarization and scaling corrections.
Determine precise cell parameters.
Collect appropriate information for an absorption correction. (Index the faces of the crystal. A highly redundant set of data is sufficient for an empirical absorption correction.)
Apply an absorption correction to the data.
(http://xrayweb.chem.ou.edu/notes/collect.html)

## Slide 8
Single crystal diffraction of X-rays
Note: The non SI unit Å is normally used.
1 Å = 10-10 m
L to K transitions produce 'Ka' emission
M to K transitions produce 'Kb' emission.
M to L transitions produce 'La' emissions.
There are several energy sublevels in the L, M, N levels so there are in fact 'Ka1' and 'Ka2' peaks which are very close to one another in energy.
Principle quantum number
n = 1  K level
n = 2  L level
n = 3  M level
etc…
Notes: 8

## Slide 9: Each element has its own characteristic x-ray spectrum
For Copper the characteristic wavelengths (λ) are:
  Cu Kα1 = 1.540Å
  Cu Kα2 = 1.544Å
  Cu K = 1.392Å
For Molybdenum they are:
  Mo Kα1 = 0.70932Å
  Mo Kα2 = 0.71354Å
  Mo K = 0.63225Å
We use MoKα (avg.) radiation
  (λ) = 0.71073Å
Or CuKα (avg.)
  (λ) = 1.54178Å
Single crystal diffraction of X-rays
A. Sarjeant
Notes: Graphically, this is what the x-ray spectrum looks like for a given element. You can see the large hump of Bremsstrahlung radiation, and the spikes for given electronic transitions.
In crystallography, for laboratory sources, we deal mostly with Copper and Molybdenum radiation.
You can see that there are two distinct peaks, and one of them is split into a doublet. The doublet is the Kα transition and it is split because there are two closely spaced energy levels in the L shell.
Looking at the x-ray wavelengths produced for each transition, you can see why these would be useful to study molecular structures. (Hint: they’re on the order of interatomic distances!).

## Slide 10
Single crystal diffraction of X-rays
http://xray0.princeton.edu/~phil/Facility/Guides/Phillips_sealed_tube.jpg
A large potential difference (ex. 50kV) is put between a tungsten filament (cathode) and a metal target (anode; ex. Molybdenum).
Electrons ejected from the filament ionize electrons from the target material. When these electrons drop back into the vacated energy levels, they give off energy partially in the form of electromagnetic radiation (and a lot of lot of heat; the tube is water cooled.)
Different metal targets emit X-rays of different wavelengths.
Beryllium windows (toxic; do not touch!) are relatively transparent to X-rays and let the X-rays escape the evacuated tube.
Notes: 10

## Slide 11
Single crystal diffraction of X-rays
Normally, X-ray lab users must become "authorized users“; these users wear badges that monitor any exposure to radiation. This is federally regulated.
Some general safety notes:
1. Know the expected path of the main X-ray beam. Always keep all parts of your body outside of this path.
2. Whenever possible, keep the safety doors to the instrument closed. For most modern instruments safeties are in place that make it impossible for the X-ray shutter to be open at the same time as the instrument doors.
3. No unauthorized personnel may defeat or override any safety features
Notes: 11

## Slide 12
Single crystal diffraction of X-rays
Some extra safety notes:
There is a serious hazard associated with possible electrical shock. The X-ray generator is a highly-regulated DC power supply that operates at an applied voltage of 50 kV, and 30-40 mA (this may vary with instrument and operator.)
The X-ray generator has several large capacitors. Even when the instrument is turned off, these capacitors store sufficient power to injure and possibly kill a person. All work on any X-ray generator should be done only by personnel trained in high-voltage electronics.
Never work above or below the generator cabinet.
Notes: 12

## Slide 13
Single crystal diffraction of X-rays
Mo X-ray tube
Lights up when shutter is open
Graphite monochromator
Collimator – Attenuates X-ray beam diameter
Beamstop (literally!)
Sample
Goiniometer
CCD Detector
Notes: 13

## Slide 14
Single crystal diffraction of X-rays
Mo or Cu Source
Monochromator
Collimator
l = 1.5418 Å
l = 0.7107 Å
Notes: 14

## Slide 15
Your structure refinement will only be as good as the data that you collect
Four things to consider:
  Your crystal
  Your instrument
  How you collect your data
  How you treat your data post-collection
“Garbage In = Garbage out” (P. Müller, 2009)

## Slide 16
Upcoming lecture on crystal growth
Earlier lecture on qualities to look for in a good crystal
Worth spending time carefully looking for the best possible crystal using a polarized microscope
Limitations:
  Crystals that desolvate readily and are not amenable to prolonged examination
  The “best” crystal may not be representative of the bulk sample.
Choosing a Crystal

## Slide 17
Normally crystals are selected to be smaller than the diameter of the beam to ensure a constant volume of irradiated matter
Crystals can be cut to size (with some practice)
Critically examine a few initial images
Crystal Mounting
https://www.bruker.com/fileadmin/user_upload/8-PDF-Docs/X-rayDiffraction_ElementalAnalysis/SC-XRD/Webinars/Bruker_AXS_Growing_Mounting_Single_Crystals_Webinar_201011026.pdf

## Slide 18
Other considerations
  Tools for mounting
  The actual mount
  Oil, epoxy, UV-curing
Data Collection Temperature
  Low temperature (ex. 100 K) to minimize thermal vibrations
  Constant temperature (even if collected close to RT, use of a low temperature device to maintain a constant temperature throughout experiment)
Crystal Mounting

## Slide 19
Crystal Mounting
https://www.bruker.com/fileadmin/user_upload/8-PDF-Docs/X-rayDiffraction_ElementalAnalysis/SC-XRD/Webinars/Bruker_AXS_Growing_Mounting_Single_Crystals_Webinar_201011026.pdf

## Slide 20: Experiment geometry
A. Sarjeant
Notes: In order to obtain a complete diffraction pattern, we need to be able to satisfy Bragg’s Law for as many sets of diffracting planes as possible. We have three choices: We can move the source to change the incident angle of diffraction, we can change the detector angle to cover all areas of space, or we can change the orientation of the crystal. Given the complicated nature of x-ray sources, it is impractical to move the source around. That leaves the sample and the detector as possible candidates.
How do we do this?
This schematic shows the geometry of a 4-cirgle goniometer. (A goniometer is any device used to measure angles and is the crux of the modern diffractometer). There are four standard angles with which crystallographers concern themselves. The line formed by the incoming x-ray beam is fixed, and the angle the detector makes with that line is the theta angle, or rather 2theta from Bragg’s Law. A second angle, coaxial with 2theta is the omega axis. Omega and 2theta rotate about an axis that is orthogonal to the direct beam. The crystal can be oriented by using two additional axes, chi and phi. The phi axis is mounted on the chi circle and both rotate with omega, as shown above.
With area detectors, the most efficient way to collect a complete set of data is by moving both the crystal and the detector. Typically, the crystal is moved in a series of steps – anywhere from 0.3deg. to 2.0deg. (we most often use 0.5deg. scans) – while the detector records reflection intensities. Because the detector cannot cover the needed resolution range in one setting, typically a series of runs is executed with the detector at different positions in 2theta, and the crystal rotating through either phi or omega.

## Slide 21: Eulerian Geometry
A. Sarjeant

## Slide 22: Kappa Geometry



2
dx
A. Sarjeant
Notes: At Northwestern, our three diffractometers have Kappa geometry, which is shown here.
It is called a Kappa geometry because instead of a Chi angle, we now have Kappa. The Kappa angle is offset by 50deg. from the omega axis. And because of this non-orthogonal axis, we are able to reach more areas of space, and put more reflections into diffracting position, than with the “Ferris wheel” chi angle. Kappa systems are a type of 4-circle instrument because they maintain four angular degrees of freedom. An older instrument that we retired last year was a type of 3-circle instrument because the chi angle was fixed and therefore we only had three angles to adjust.
Our Kappa goniometers also have movable detectors which help accommodate samples with widely disparate unit cell sizes as well as the swinging kappa arm.
The cold stream blows a stream of nitrogen gas over the sample keeping it cold, keeping it protected from water and oxygen in the atmosphere, and keeping it from decomposing or losing solvent. X-ray diffraction experiments will determine atomic positions. At lower temperatures, the thermal motions of the atoms are much smaller, making it much easier to get accurate atomic positions. Yet another reason to keep the sample cold while you are collecting data.
The direct beam passes through the collimator from the source tube to the crystal. The crystal is mounted on the end of a fiber on a goniometer head. The goniometer head can be translated in three dimensions to ensure the crystal is in the center of the x-ray beam and not precessing as it rotates. To make centering easier, a video camera is mounted on the diffractometer base, focused on the crystal.
The beamstop is situated behind the crystal, between the source and the detector to catch the primary beam. Most of the x-ray radiation will pass through or around the crystal (if it is smaller than the beam). If the direct beam is allowed to hit the detector, it will overload the electronics and wash out all the diffracted beams and you won’t see any spots. [Think about taking a picture of a back-lit subject. Typically the light overwhelms everything and you can’t see the details of your subject.]
This diffractometer also has a movable detector which helps accommodate different unit cell sizes as well as the swinging kappa arm.

## Slide 23
A. Sarjeant
Notes: This is an animation of the movement of a Kappa goniometer. Thanks, Wikipedia!

## Slide 24
Single crystal diffraction of X-rays
Recall: The diffraction pattern does not depend on translation, but does rotate if the lattice is rotated.
The following video shows the images from an X-ray diffraction data collection:
http://ruppweb.org/data/vta1.wmv
Notes: 24

## Slide 25
Regular maintenance
Correctly aligned
How do you know?
  Stable test crystal that is regularly collected, with comparison to previous results.
  When in doubt about your own instrument, recollect the test crystal.
Instrumental Optimization

## Slide 26
Reflection intensities are generally weaker at higher resolutions, but high angle data contains important structural information.
IUCr generally
recommends a
minimum resolution of
0.84 Å.
(How does this
relate to Bragg’s Law?)
Data Collection Strategy: Maximum Resolution

## Slide 27
The normal range of X-H bonds is ~0.80-0.95 A. At 53o these separations can be resolved.
Problem: The Acta Cryst standard for 2q collections is a minimum cut-off of 53o. Why do you think that is?
Solution: Employing Bragg’s law with l = 0.7107 Å (Mo-Ka radiation) and q = 26.5o:
The normal range of X-H bonds is ~0.80-0.95 A. At 53o these separations can be resolved.
Data Collection Strategy: Maximum Resolution

## Slide 28
Bragg’s Equation
2q = 17o (l = 0.7107 Å)
Old protein structures
No distinct atomic positions can be identified
2q = 41.6o (l = 0.7107 Å)
Small molecule solution possible.
Refinement of atomic positions will have large associated errors.
2q = 50.7o (l = 0.7107 Å)
See previous example
This should lead to a publishable result.
Reprinted from Interpretation of Crystal Structure Determinations. Copyright 2005 Huub Jooijman, Bijvoet Center for Biomolecular Research and Structural Chemistry, Utrecht Univeristy.
CH495 Dr. L. Dawe Fall 2014
Notes: 28

## Slide 29
Data completeness is the data actually collected compared to what is the unique data for the given crystal symmetry.
Software will allow you to determine a data collection strategy to yield 100% completeness.
Some crystallographers have developed their own collection strategies (based on presumed low symmetry and experience.)
Data Collection Strategy: Data Completeness

## Slide 30
Data completeness is the data actually collected compared to what is the unique data for the given crystal symmetry.
Software will allow you to determine a data collection strategy to yield 100% completeness.
Some crystallographers have developed their own collection strategies (based on presumed low symmetry and experience.)
Data Collection Strategy: Data Completeness

## Slide 31
Data Collection Strategy: I/s
Average measured intensity/estimated noise
Ideally should be as
high as possible (~10
throughout the data
set)
Values less than 2
are essentially noise
Decisions about where
to cut off your resolution?

## Slide 35
Two-fold symmetry about b-axis

## Slide 36
Two-fold symmetry about b-axis
Mirror symmetry perp to b-axis

## Slide 40
Multiplicity of Observation (MoO) refers to multiple measurements of the same, or symmetry equivalent, reflection, obtained from a different crystal orientation.
Higher values of MoO should yield better statistics
Higher symmetry crystals require less images to obtain equivalent MoO to lower symmetry crystals
One approach is to collect all crystals as though they were triclinic (over-estimating symmetry can yield incomplete data.)
Data Collection Strategy: Multiplicity of Observations

## Slide 41
Chart: Residuals vs. multiplicity
| 1 | 2 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19
 | 4.02 | 3.84 | 3.68 | 3.44 | 3.28 | 3.15 | 3.09 | 3.06 | 2.98 | 2.93 | 2.92
 | 8.98 | 8.85 | 8.75 | 8.36 | 8.06 | 7.86 | 7.79 | 7.65 | 7.52 | 7.45 | 7.45
wR2 = 0.074
R1 = 0.029

## Slide 42
Chart: Avg C-C bond precision
| 1 | 2 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19
 | 0.003 | 0.003 | 0.003 | 0.0026 | 0.0023 | 0.0022 | 0.0022 | 0.0022 | 0.0021 | 0.0021 | 0.0021

## Slide 43
Modifications to measured I(hkl) are required to correct for geometry of measurement
Essential to yield high quality accurate data for solution and refinement.
Some correction factors include:
  Lorentz factor (accounts for time required for a Bragg reflection to cross the surface of the sphere of reflection)
  Polarization factor (polarization of the incident X-ray beam)
  Absorption (intensity of measure reflections is reduced by the absorption of X-rays by the crystal)
Processing

## Slide 44
Processing: Corrections
For a small crystal completely bathed in a uniform beam of radiation, the integrated intensity, I, is given by:
I = Io (re)2 (Lp/A) (λ/Ω) (F/V)2 λ2υ
The quantity re = e2/mc2 = 2.82 × 10-13 cm is the classical radius of an electron. V is the unit cell volume; υ is the volume of the crystal. Ω is the angular velocity of the sample as the peak moves through the Ewald sphere. Correction terms include the Lorentz correction, L, the polarization correction, p, and the absorption correction, A.
http://xrayweb.chem.ou.edu/notes/collect.html#correction

## Slide 45
The absorption of X rays follows Beer's Law:
I / Io = exp(-μ × t)
where I = transmitted intensity, Io = incident intensity, t = thickness of material, μ = linear absorption coefficient of the material. The linear absorption coefficient depends on the composition of the substance, its density, and the wavelength of the radiation. Since μ depends on the density of the absorbing material, it is usually tabulated as the related function mass absorption coefficient μm = (μ / ρ).
The linear absorption coefficient is then calculated from the formula:
μ = ρ ∑ (Pn / 100) × (μ / ρ) = ρ ∑ (Pn / 100) × μm
where the summation is carried out over the n atom types in the cell, and Pn is the percent by mass of the given atom type in the cell.
http://xrayweb.chem.ou.edu/notes/collect.html#corrections
Processing: Absorption Corrections

## Slide 46
Crystals were ground or cut to be approximately spherical in order to minimize unequal absorption effects
Numerical absorption corrections require accurate information about crystal shape by way of indexing a crystal’s faces. Analytical absorption corrections are accomplished by mathematically dividing the sample into very small pieces and calculating the transmittance for each piece of the crystal for each reflection measured. This can be difficult for crystals with many closely spaced faces. The hkl indices of faces and their distances from the center of the crystal are required. Less common now, but still used for very strongly absorbing materials and charge density studies.
Modern semi-empirical methods are based on measurement of equivalent reflections and work well when there is a high multiplicity of observations. By comparing the intensities from the redundant measurements, an absorption surface for the sample is calculated.
http://xrayweb.chem.ou.edu/notes/collect.html#corrections
Processing: Absorption Corrections

## Slide 47
Symmetry-equivalent intensity data are merged using the following relationship
F2 = ∑ ωj Fj2 / ∑ ωj
where the summations are over the set of symmetry-equivalent data. In this formula, weights can be either from statistics (ω = 1/σ(F2)) or unit values . (http://xrayweb.chem.ou.edu/notes/collect.html#merge)
When comparing unique data to total data collected, there are a variety of residuals that can be calculated as a measure of internal data consistency
For example: _diffrn_reflns_av_R_equivalents with is the residual for symmetry-equivalent reflections used to calculate the average intensity.
Lower merging R-factors indicate better datasets (would like to see less than 10% over the entire range of resolution)
Data Collection Strategy: Merging Residuals

## Slide 48
Where can I find this info?

## Slide 49
A complete collection of at least one crystal structure for all of the 230 space groups:
https://crystalsymmetry.wordpress.com/230-2/
Something fun and marginally related